## Instructions
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1
Latest Template Revision: | 2017-07-07
Tab Name: | Instructions
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Reference Test Procedure
10 CFR 430 Subpart B Appendix E: Uniform Test Method for Measuring the Energy Consumption of Water Heaters
Table of Contents
Tab | Contents
Instructions | Instructions and table of contents
General Info & Test Results | Lab information, product information and test results
Setup & Instrumentation | Instrumentation requirements and space for sensor placement descriptions
Photos | Inputs for photographs
Test Conditions | Input for test condition measurements
Max GPM Test | Measurement inputs and automated calculations
24 Hr test | Measurement inputs and automated calculations
Comments | Inputs for report template user to provide comments
Report Sign-Off Block | Report review history
Drop-Downs | Drop-downs used
Version Control | Revision history
LEGEND
Tabs | Tabs with input cells
Cells | Input cell
Auto-populated cell
Provided data
NOT USED
Instructions for Completing this Template
Follow the steps below, filling in all input cells (shaded light blue) in each tab you are instructed to complete. Using TAB to "hop" from input cell to input cell is useful, but does not ensure that all input cells are reached. To guarantee that you enter all required information, you must visually scan for light blue cells in the entire area bounded by yellow-shaded cells. This test report template is only for tankless electric water heaters, and is not applicable for the other water heater types listed in the test procedure.
Important: Start with a clean (unused) template copy for each new report. Enter only data and information that are unique to the unit tested and the current test of that unit. All abbreviations and variable names should be consistent with the reference test procedure.
STEP: | FILL IN INPUT CELLS IN THE CORRESPONDING TAB:
Step 1 | General Info & Test Results
Step 2 | Setup & Instrumentation
Step 3 | Photos
Step 4 | Test Conditions
Step 5 | Max GPM Test
Step 6 | 24 Hr Test
Step 7 | Comments
Step 8 | Report Sign-Off Block

## General Info & Test Results
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | General Info & Test Results
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Lab Information |  | Test Results
Lab Name: |  | Variable | Measurement | Units
Lab Location: |  | Maximum GPM Rating |  | GPM @ 77°F Rise
|  | Energy Factor, Ef |  | -
Test Information |  | Annual Energy Consumption, Eannual |  | KWH/Yr
Date Test Started: | [MM/DD/YYYY]
Date Test Finished: | [MM/DD/YYYY] | NOTE: This is only a copy; sign off is done in the Report Sign-Off Block tab
Test Report Sign-Off Block
Product Information |  | We certify that the information and data in this report: (1) were obtained from the specific test unit under test; (2) were obtained during the specific test being reported; (3) were not copied from any other source, except where instructed to do so; and (4) were not altered or modified in any way.
Manufacturer:
Brand:
Manufacturer Model Number: |  | Role |  | Date | Entity
Serial Number: |  | Test Completion |  | [MM/DD/YYYY] | [Test Lab Name]
Date Manufactured: |  | Template Completion |  | [MM/DD/YYYY] | [Test Lab Name]
Date Product Received: | [MM/DD/YYYY] | Report Review by Test Lab |  | [MM/DD/YYYY] | [Test Lab Name]
Condition as Received: |  | Report Review by Test Lab |  | [MM/DD/YYYY] | [Test Lab Name]

## Setup & Instrumentation
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 |  | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | Setup & Instrumentation
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Setup (this table should include instrumentation, sensors, and all equipment used during testing)
Instrument Type | Brand | Model # | Sensor Location | Accuracy | Date of Last Calibration | Deadline for Next Calibration
Please describe the test setup:
Describe placement of sensor used to measure ambient temperature:
Describe placement of sensors used to measure supply water temperature:
Describe placement of sensors used to measure supply water pressure:

## Photos
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | Photos
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
1. Nameplate showing model number and serial number (if applicable) |  |  | 2. FTC EnergyGuide label (if present)
3. Photos of test unit from all sides (including photo of control panel, if applicable)
4. Overall view of unit installed |  |  | 5. Exact placement of all visible sensors on, in, or around the device
6. View of water inlet and outlet plumbing (including thermocouples and any insulation) and vent pipes and terminal |  |  | 7. Additional photos (if necessary)

## Test Conditions
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1
Latest Template Revision: | 2017-07-07 |  | Back to Instructions tab
Tab Name: | Test Conditions
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Test Conditions
Test Start Date: | [MM/DD/YYYY]
1. Ambient Temperature (in accordance with Section 2.2 of Appendix E).
The ambient air temperature shall be maintained between 65.0 °F and 70.0 °F (18.3 °C and 21.1 °C) on a continuous basis.
Amb. Temperature (°F)
| Min | Max
Max GPM Test
24-Hr Simulated Use Test
2. Supply Water Temperature (in accordance with Section 2.3 of Appendix E).
The temperature of the water being supplied to the water heater shall be maintained at 58 °F ±2°F (14.4 °C ±1.1 °C) throughout the test.
Supply Water Temperature (°F)
| Min | Max
Max GPM Test
24-Hr Simulated Use Test
3. Supply Water Pressure (in accordance with Section 2.5 of Appendix E).
During the test when water is not being withdrawn, the supply pressure shall be maintained between 40 psig (275 kPa) and the maximum allowable pressure specified by the water heater manufacturer.
Supply Water Pressure (psig)
| Min | Max
Max GPM Test
24-Hr Simulated Use Test
4. Power Input Determination (in accordance with Section 2.6.1 of Appendix E).
Maintain the electrical supply voltage to within 1% of the center of the voltage range sepcified by the water heater and/or heat pump manufacturer.
Electrical Supply Voltage (V)
Manufacturer Recommended Voltage
Min Voltage (V)
Max Voltage (V)
Within 1% of voltage range? | Yes

## Max GPM Test
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1
Latest Template Revision: | 2017-07-07 |  | Back to Instructions tab
Tab Name: | Max GPM Test
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
10 minute Max Flow Rate Test (in accordance with Sections 5.2.1 to 5.2.3 of Appendix E)
Test Start Date: | [MM/DD/YYYY]
Select Input Control:
From Section 5.2.1:
Determine the flow rate (Vmax) at which the WH can discharge water at 135±5 °F (57.2 °C ±2.8 °C), when set to the maximum firing rate.
Vmax |  | GPM
From Section 5.2.2 (For Variable Input Water Heaters):
If the water heater incorporates a controller that permits operation at a reduced input rate, adjust the flow rate as necessary to achieve a discharge water temperature of 135 °F ±5 °F (57.2 °C ±2.8 °C) while maintaining the minimum input rate. Record the corresponding flow rate as Vmin . If an outlet temperature of 135 °F ±5 °F (57.2 °C ±2.8 °C) cannot be achieved at the minimum flow rate permitted by the water heater, record the flow rate as Vmin and the corresponding outlet temperature as Tmin .
Vmin |  | GPM
Tmin |  | °F
From Section 5.2.3:
Establish normal water heater operation at the full input rate with the discharge water temperature set in accordance with Section 5.2.1. Either collect the withdrawn water for later measurement of the total mass removed, or alternatively, use a water meter to directly measure the water volume removed.
Select mass or volume basis for data acquisition: |  | (This selection will not affect final results.)
Test Data (10 Minute Run)
Mass of Water (M10m) |  | lbs, if scale is used
Volume of Water (V10m) |  | gallons, if flow meter is used
Average Delivery Temp. |  | °F
Average Inlet Temp. |  | °F
Density of water based on average delivery temp. |  | lbs/gal
| Calculated Result (unrounded) | Calculated Result (rounded per significant figures conventions and instrumentation resolution)
Maximum GPM (Fmax) = |  |  | GPM @ 77°F Rise
Temperature Data
After recording the scale or water meter reading, initiate water flow throughout the water heater, record the inlet and outlet water temperatures beginning 15 seconds after the start of the test and at subsequent 5-second intervals throughout the duration of the test.
Time (secs) | Temperature (°F) |  | Total Mass
| Inlet Temperature (°F) | Delivery Temperature (°F) | lbs
15
20
25
30
35
40
45
50
55
60
65
70
75
80
85
90
95
100
105
110
115
120
125
130
135
140
145
150
155
160
165
170
175
180
185
190
195
200
205
210
215
220
225
230
235
240
245
250
255
260
265
270
275
280
285
290
295
300
305
310
315
320
325
330
335
340
345
350
355
360
365
370
375
380
385
390
395
400
405
410
415
420
425
430
435
440
445
450
455
460
465
470
475
480
485
490
495
500
505
510
515
520
525
530
535
540
545
550
555
560
565
570
575
580
585
590
595
600 |  |  |  | Total

## 24 Hr Test
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 |  | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | 24 Hr Test
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
24 Hour Test Measurements (according to section 5.2.4 of Appendix E)
Test Start Date: | [MM/DD/YYYY]
Select mass or volume basis for data acquisition: |  |  | (This selection will not affect final results.)
| Draw #1 | Draw #2 | Draw #3 | Draw #4 | Draw #5 | Draw #6
Target Flow Rate [GPM] | 0 | 0 | 0 | 0 | 0 | 0 | (Draws 4-6 will vary depending on fixed/variable input.)
Target for Gallons Used | 10.7167 | 10.7167 | 10.7167 | 10.7167 | 10.7167 | 10.7167 | (Draws 4-6 will vary depending on fixed/variable input and Tmin.)
Tolerance Allowed [Gallons] | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5
Water Drawn [Gallons] | [6 empty] | (Select Mass or Volume Basis above.)
Avg. Water Temp In, Tin [°F] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | (See data acquisition table below.)
Avg. Water Temp Out, Tdel [°F] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | (See data acquisition table below.)
Density of Water at point where Vol is measured, ρ [lb/gal] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | (Based on inlet temperatures.)
Mass Water Removed, M [lb] | Basis Unselected | Basis Unselected | Basis Unselected | Basis Unselected | Basis Unselected | Basis Unselected
Volume of Water Removed, V [gal] | Basis Unselected | Basis Unselected | Basis Unselected | Basis Unselected | Basis Unselected | Basis Unselected
Specific Heat of Withdrawn Water, Cp [Btu/lb-F] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | (Based on average of inlet and outlet temperatures.)
Specific Heat of Withdrawn Water, Cp [Btu/lb-F] at (58F+135F)/2 | 0.99798 | 0.99798 | 0.99798 | 0.99798 | 0.99798 | 0.99798
Temperature Rise [°F]: | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Watt-Hrs Used:
Electric BTU: | 0 | 0 | 0 | 0 | 0 | 0
Mi Cpi (Tdel,i - Tin,i) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Mi Cpi (135F-58F) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Mi Cp@(58F+135F)/2 (135F-58F) | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Watt-Hrs used between End of 6th Draw and End of 24-hr Period [Watt-Hrs]:
Electricity Consumption between End of 6th Draw and End of 24-hr Period [Btu] | 0
Electric Meter Reading between End of 6th Draw and End of 24-hr Period [Watt-Hrs]:
Total Watt-Hrs Used over 24 Hr Test:
Total Electric BTU Used over 24 Hr Test: | 0
ηr,max | #VALUE! |  | ηr,min | #VALUE! | (ηr,max and ηr,min will be the same for fixed input WH.)
ηr | #VALUE!
Modified Daily Water Heating Energy Consumption (Qdm)
Qd | 0 | BTU
QHW | #VALUE! | BTU
QHW,77 | #VALUE! | BTU
Qdm | #VALUE! | BTU
| Calculated Result (unrounded) | Calculated Result (rounded per reference test procedure)
Ef | #VALUE! | #VALUE! | -
Average Electricity Cost = |  | $/KWH
| Calculated Result (unrounded) | Calculated Result (rounded per significant figures conventions and instrumentation resolution)
Eannual | #VALUE! |  | KWH/Yr
Temperature Data
Complete each draw to meet target for gallons used.
Time (sec) | Draw 1 |  | Draw 2 |  | Draw 3 |  | Draw 4 |  | Draw 5 |  | Draw 6
| Tin [°F] | Tdel [°F] | Tin [°F] | Tdel [°F] | Tin [°F] | Tdel [°F] | Tin [°F] | Tdel [°F] | Tin [°F] | Tdel [°F] | Tin [°F] | Tdel [°F]
15
20
25
30
35
40
45
50
55
60
65
70
75
80
85
90
95
100
105
110
115
120
125
130
135
140
145
150
155
160
165
170
175
180
185
190
195
200
205
210
215
220
225
230
235
240
245
250
255
260
265
270
275
280
285
290
295
300
305
310
315
320
325
330
335
340
345
350
355
360
365
370
375
380
385
390
395
400
405
410
415
420
425
430
435
440
445
450
455
460
465
470
475
480
485
490
495
500
505
510
515
520
525
530
535
540
545
550
555
560
565
570
575
580
585
590
595
600
605
610
615
620
625
630
635
640
645
650
655
660
665
670
675
680
685
690
695
700
705
710
715
720
725
730
735
740
745
750
755
760
765
770
775
780
785
790
795
800
805
810
815
820
825
830
835
840
845
850
855
860
865
870
875
880
885
890
895
900
905
910
915
920
925
930
935
940
945
950
955
960
965
970
975
980
985
990
995
1000
1005
1010
1015
1020
1025
1030
1035
1040
1045
1050
1055
1060
1065
1070
1075
1080
1085
1090
1095
1100
1105
1110
1115
1120
1125
1130
1135
1140
1145
1150
1155
1160
1165
1170
1175
1180
1185
1190
1195
1200
1205
1210
1215
1220
1225
1230
1235
1240
1245
1250
1255
1260
1265
1270
1275
1280
1285
1290
1295
1300
1305
1310
1315
1320
1325
1330
1335
1340
1345
1350
1355
1360
1365
1370
1375
1380
1385
1390
1395
1400
1405
1410
1415
1420
1425
1430
1435
1440
1445
1450
1455
1460
1465
1470
1475
1480
1485
1490
1495
1500
1505
1510
1515
1520
1525
1530
1535
1540
1545
1550
1555
1560
1565
1570
1575
1580
1585
1590
1595
1600
1605
1610
1615
1620
1625
1630
1635
1640
1645
1650
1655
1660
1665
1670
1675
1680
1685
1690
1695
1700
1705
1710
1715
1720
1725
1730
1735
1740
1745
1750
1755
1760
1765
1770
1775
1780
1785
1790
1795
1800
1805
1810
1815
1820
1825
1830
1835
1840
1845
1850
1855
1860
1865
1870
1875
1880
1885
1890
1895
1900
1905
1910
1915
1920
1925
1930
1935
1940
1945
1950
1955
1960
1965
1970
1975
1980
1985
1990
1995
2000
2005
2010
2015
2020
2025
2030
2035
2040
2045
2050
2055
2060
2065
2070
2075
2080
2085
2090
2095
2100
2105
2110
2115
2120
2125
2130
2135
2140
2145
2150
2155
2160
2165
2170
2175
2180
2185
2190
2195
2200
2205
2210
2215
2220
2225
2230
2235
2240
2245
2250
2255
2260
2265
2270
2275
2280
2285
2290
2295
2300
2305
2310
2315
2320
2325
2330
2335
2340
2345
2350
2355
2360
2365
2370
2375
2380
2385
2390
2395
2400
2405
2410
2415
2420
2425
2430
2435
2440
2445
2450
2455
2460
2465
2470
2475
2480
2485
2490
2495
2500
2505
2510
2515
2520
2525
2530
2535
2540
2545
2550
2555
2560
2565
2570
2575
2580
2585
2590
2595
2600
2605
2610
2615
2620
2625
2630
2635
2640
2645
2650
2655
2660
2665
2670
2675
2680
2685
2690
2695
2700
2705
2710
2715
2720
2725
2730
2735
2740
2745
2750
2755
2760
2765
2770
2775
2780
2785
2790
2795
2800
2805
2810
2815
2820
2825
2830
2835
2840
2845
2850
2855
2860
2865
2870
2875
2880
2885
2890
2895
2900
2905
2910
2915
2920
2925
2930
2935
2940
2945
2950
2955
2960
2965
2970
2975
2980
2985
2990
2995
3000

## Comments
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | Comments
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Comments

## Report Sign-Off Block
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 |  | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | Report Sign-Off Block
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Test Report Sign-Off Block
By signing in the space below, we certify that the information and data in this report: (1) were obtained from the specific test unit under test; (2) were obtained during the specific test being reported; (3) were not copied from any other source, except where instructed to do so; and (4) were not altered or modified in any way.
Role |  | Date | Entity
Test Completion |  | [MM/DD/YYYY] | [Test Lab Name]
Template Completion |  | [MM/DD/YYYY] | [Test Lab Name]
Report Review by Test Lab |  | [MM/DD/YYYY] | [Test Lab Name]
Report Review by Test Lab |  | [MM/DD/YYYY] | [Test Lab Name]

## Drop-Downs
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 |  | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | Drop-Downs
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Input Control Type |  | Basis |  | Y_N
Fixed |  | Mass |  | Yes
Variable |  | Volume |  | No

## Version Control
Title Block
Test Report Template Name: | Water Heater, Tankless Electric
Version Number: | v2.1 | Back to Instructions tab
Latest Template Revision: | 2017-07-07
Tab Name: | Version Control
File Name: | Water Heater, Tankless Electric - v2.1.xlsx
Test Completion Date: | [MM/DD/YYYY]
Revisions List
Version | Date
1.0 [draft] | 2013-03-14
1.0_draft1 | 2013-05-29
1.0_draft2 | 2013-05-31
1.0_draft3 | 2013-07-06
1.0_draft4 | 2013-07-08
1.0_draft5 | 2013-07-23
1.0_draft6 | 2013-09-06
1.0_draft7 | 2013-09-11
1.0_draft8 | 2014-06-25
v1.0 | 2014-10-08
v2.0 | 2015-06-05
v2.1 | 2017-07-07